## Slide 1: Targeted Proteomics in Skyline: Experimental Design, Method Development, and Statistical Analysis
Northeastern University
May 1-3, 2017
Brendan MacLean, Lindsay Pino, Susan Abbatiello

## Slide 2: Goals and Concepts Day 1
Set up a targeted LC-MS/MS method
  Empirical data vs. information from literature
Skyline
Refine and optimize method
  Peptide selection and transition refinement
  MS settings and LC conditions
  iRT

## Slide 3: Goals and Concepts Day 2
Absolute Quantitation and Method Validation
  Appropriate experimental design
  Internal standards
Calibration curves
  Appropriate design and execution
  Statistical analysis, LOD/LOQ calculations
Peptide Stability
Method Reproducibility
Method Repeatibility

## Slide 4: Goals and Concepts Day 3
System Suitability
  Method design
  Analyte selection
Data evaluation
AutoQC

## Slide 5: Instructors
Brendan MacLean
Lindsay Pino
Susan Abbatiello
D. R. Mani
Cyril Galitzine

## Slide 6: Designing a Targeted LC-MS/MS Method for High Confidence Peptide Quantitation
Susan Abbatiello
Northeastern University Skyline Short Course
May 1, 2017

## Slide 7: Current Application of LC-MS for Discovery of Biomarkers
Adapted from Rifai, Gillette and Carr Nat. Biotech.2006
Discovery
Tissue
Proximal fluids
Clinical
Validation
Blood
Population
Verification
Blood
Population
Bio-Specimens
Plasma
Tissue
Proximal fluids
Found in blood?
higher in cancer?
Biomarkers worth
evaluating
biomarker
candidates
“hypotheses”
untargeted proteomics
genomics
> 10,000
Analytes
4 - 10
Analytes
10’s
Samples
1000’s
Samples
100’s
Samples
100’s
Analytes

## Slide 8: Overview
Considerations for Designing Targeted Experiments
Quantitation by MS?!?!
Mass Analyzers
Evaluating your data

## Slide 9: What are your goals?
Precision/Reproducibility
Limit of Detection (LOD)
Present
Not Detected
Detection
Quantitation
- Or -
Accuracy
Precision/Reproducibility
Linearity
Lower Limit of Quantitation (LLOQ)
Upper Limit of Quantitation (ULOQ)

## Slide 10: Considerations in Developing Targeted Assay
Why was protein candidate selected?
  Empirical evidence (and is any of it MS data)?
  Biological significance and biological variation?
What amount of protein might be in sample?
  Is enrichment necessary?
  How much sample is available?
How many targets?
  How many transitions per peptide?
Cost
  IS peptides cost ~$500 (>1000’s experiments possible)
  Method development time
  MS and LC
Time
  Sample acquisition time
  Data analysis time
Biological interferences
  How unique is the protein or peptide?
  Potential for false positives

## Slide 11: Never Underestimate Literature
Search for your proteins of interest
  BLAST sequences, double-check
Can you get ≥100 amol onto a column?
How much sample will you need?
  10 L of blood?
  10 mg cell lysate?
Investigate if the protein has been detected and what the sample prep was
Is the site-specific modification more important? Can you locate it by MS/MS?
Sometimes things haven’t been published, not because they haven’t been thought of or tested, but because there were no positive results (or it’s just impossible!!!).

## Slide 12
“Interesting” Biomarkers are at Low Concentrations and Masked by More Abundant Proteins
Slide courtesy of N. Leigh Anderson

## Slide 13: Sample Enrichment and Nanoflow Chromatography Improve Sensitivity in LC-MS
Sample Enrichment (Plasma) Equivalent Sample Volume
Plasma Depletion (protein level) ~220 nL per inj.
Orthogonal Chromatography (peptide level) ~500 - 1500 nL per inj.
Antibody Enrichment (peptide level) ~100-1000 uL per inj.
1 ug neat plasma (no enrichment)
~ 15 nL
Protein Sample Cell Lysate Tissue Plasma
Protein-Level
Immunoaffinity Cellular Fractionation
Peptide-Level
IMAC (Phospho)
Immunoaffinity
Fractionation
Quantitative
MRM-MS Method
10-100 uL plasma
107 cells
Keshishian et al. (2007) Mol Cell Proteomics; Whiteaker et al. (2007) Anal Biochem
100’s amol per 1 ug protein (no enrichment)
100’s amol per 1 mg protein (IMAC enrichment)
Digest to peptides
nLC
*
*

## Slide 14: Chemical Noise Interferences: Balancing Signal vs Noise
Modes of separation
2: LC/MS, GC/MS, MS/MS
3: LC/MS/MS
4: Mudpit (LC/LC/MS2)
4: LC/Ion Mobility/MS/MS
Selectivity

## Slide 15: Chromatographic Flow Rate Affects Sensitivity
75µm ID
150µm ID
300µm ID
2.1mm ID
1mm ID
75µm ID
150µm ID
300µm ID
2.1mm ID
1mm ID
Slide courtesy of Jim Murphy, Waters Corporation

## Slide 16: Quantitation by MS?!?
Is it even possible….

## Slide 17
Quantitation by Mass Spectrometry
…a History
1972: Moore, L.J. and L.A. Machlan. High accuracy determination of calcium in blood serum by isotope dilution mass spectrometry. Analytical Chemistry 44:2291-2296.
1996: Barr, J.R., et al. Isotope dilution-mass spectrometric quantification of specific proteins: model application with apolipoprotein A-1. Clinical Chemistry 42:1676-1682.
1999: Gygi, S.P. et al. Quantitative analysis of complex protein mixtures using isotope-coded affinity tags. Nature Biotechnology 10:994-999.
2002: Ong, S.E, et al. Stable isotope labeling by amino acids in cell culture. SILAC, as a simple and accurate approach to expression proteomics. Molecular & Cellular Proteomics 1: 376-386
2003: Gerber, S.A., et al. 2003. Absolute quantification of proteins and phosphoproteins from cell lysates by tandem MS. PNAS 100:6490-6495.

## Slide 18: Quantitation (remember Beer’s Law?)
S = k*c + b
S = signal (MS intensity) k = (ionization) coefficient c = concentration b = background signal
A = εbc
A = signal (absorbance) ε = extinction coefficient c = concentration b = path length
??

## Slide 19
Q: Why does this work?
A: Internal Standard shares all characteristics of analyte
– EXCEPT MASS!
Quantitation by Stable Isotope Dilution MS
13C615N4 Arginine, D 10 Da
13C615N2 Lysine, D 8 Da

## Slide 20: Peptide Fragmentation
Reproducible fragments can be predicted and calculated, based on peptide sequence of amino acids
http://www.alchemistmatt.com/mwthelp/peptidefragmodelling.htm

## Slide 21: SID-MS/MS Peptide Example

## Slide 22: SID-MS Experimental Design
Time (min)
0
20
40
60
80
100
D S
Protein mixture
Digest
Peptides
Desalt
Extract Data
Peak area analyte
Peak area IS
Peak area = ratio

## Slide 23: SID-MS Response
Spike fixed amount of heavy-isotope peptide
Vary amount of light-isotope peptide
Plot Ratio of MS response vs Ratio of Amount
y = 1.0x + b
Model Experiment

## Slide 24: What do you want to Target?
Empirical LC-MS/MS data
  DDA runs from “discovery” samples
  Pure protein standards
  Synthetic peptides
Protein targets from literature
  Potential peptides previously studied by MS
  Proteins in pathways of interest
No idea?

## Slide 25: Empirical LC-MS/MS Data
Consult the data!
Empirical data are your best starting point
Peptides fragment similarly on different instruments, so use HRAM data as your guide!

## Slide 26: How do I use empirical data from LC-MS/MS experiments?
Search the data against proteomic libraries
Identify peptides with certain confidence
Generate a list of peptides unique to a protein
Target those peptides!

## Slide 27: Protein Targets from Literature
“Hypothesis-driven” vs “Discovery-driven”
  Data other than MS data point to certain targets
  Proteins or protein pathways are indicated
  No peptide data
http://www.bio-rad.com

## Slide 28: Protein Targets from the Internet!
Consult the data!
  Check publicly available databases for empirical results
  HPM
  PRIDE
  GPMDB
  PeptideAtlas
  SRMAtlas
  NISTLibraries
  Panorama
  Passport
“A month in the laboratory can often save an hour in the library” – Frank Westheimer, Harvard Chemist

## Slide 29: Protein (peptide) Targets from Public Resources: Passport (http://passport.maccosslab.org/)

## Slide 30: Protein (peptide) Targets from Public Resources: CPTAC Assay Portal (https://assays.cancer.gov)
Assays developed with internal standards
Experimental details
LOD/LOQ
Precision
Repeatibility
https://assays.cancer.gov/CPTAC-795

## Slide 31: No Idea?
Generate the Data!
Start with general “rules”
  Proteotypic peptide
  Missed cleavages
  Optimal Length
  Not close to N-terminus
  Exclude labile amino acids
    Met (oxidation)
    N-terminal Gln (cyclization)
    Cysteine (reduction/alkylation)
  Beware of PTMs

## Slide 32: Examples Using Empirical Data
Q: I have spectral libraries! How do I pick the best peptide transitions?
A: Look at the data!
Selection of 123 target peptides
Top 8-10 Product Ions
Selection of best 3-5 ions
CE Calculation
Refined Transition List for
Method Optimization
DDA LC-MS/MS, Database search
Spectral Library
SRM or PRM

## Slide 33: Which charge state do you pick?

## Slide 34: Some peptides are beautiful…

## Slide 35: Some peptides are not…

## Slide 36: For best results…
Explore the data to understand it.
Use predictions and algorithms as a guide
  Check the data manually…you might learn something!
  You’ll build experience in generating targeted methods!
  You’ll become more confident in your data!
Credit to: Various internet searches for elephants, Mr. Snuffleupagus and chameleons

## Slide 37: Choices of Mass Spectrometers
Triple Quadrupole or High Res?
(do you have a choice?)
http://people.stfx.ca/tsmithpa/chem361/labs/ms.html

## Slide 38: Differences in Mass Analyzers
Low Resolving Power
High Resolving Power
https://www.intechopen.com/source/html/44881/media/image2.png
QqQ
Ion Trap
FTICR and Orbitrap
Quadrupole-Time of Flight

## Slide 39: What MS Instrument to Use?
HRAM
High Resolution
Accurate Mass
Generates all product ions
No CE optimization needed
Precursors can be monitored
Higher Cost
20-100 spectra/sec
  ~100 peptides
QqQ
Low resolution
Nominal mass
Detects product ions on a list
CE optimization possible
Precursors not usually monitored
Lower Cost
200-500 SRMS/sec
  ~100 peptides (5 SRMS each)

## Slide 40: Instrument Cycle Time*
The amount of time it takes to acquire all data points in a given timeframe
Choose based on your chromatography
Aim for 10-15 points across the chromatographic peak
  Dwell time (plus overhead/interscan delay) for QqQs
  Injection time and detection time for HR
Current QqQs can target ~500 transitions/second (~100 peptides)
Current HR-MS can target 10-50 peptides/sec
Dwell time + “overhead” or “interscan delay”
*Not to be confused with “duty cycle”, which is the percentage of time your instrument is actively producing data.

## Slide 41: Where does the time go?
QqQ
  Interscan delay: Time to clear quads for next experiment (1-5 msec/transition)
  Dwell Time: Time spent monitoring signal from each transition (≥ 1msec/transition)
PREDICTABLE
  Independent of ion flux
HRAM
  Injection time: Max time to accumulate ions in trap
  Detection time: frequency at which data can be detected/transformed (depends on Resolving Power)
NOT PREDICTABLE
  Dependent upon ion flux

## Slide 42: SRM vs PRM vs IS-PRM
Detection is improved by targeted detection during PRM
Ions can be accumulated, offering more sensitivity than SRM
Targeting is real-time, no need for manual scheduling
Gallien, et al, MCP 2015 e-pub

## Slide 43: Looking at your Data: Quickly Screen Data in Skyline
Notes: Idea for rapidly screening large data sets with either many sample files or many peptides.
Panels show a response curve for a peptide (Red) with a heavy-labeled peptide spiked in as the standard (Blue).
Here, RT plots are shown in the upper right and the peak area CV plots are in the lower right.

## Slide 44: Interference Visualization
Heavy Peptide Transitions
Light Peptide Transitions
Notes: When the view of the Peak Area plots is changed so that the signal is normalized to 100%, you can quickly observe for interferences.
Here, the heavy peptide has the same signal in all samples, the brown transition having the largest contribution to the overall signal.
When looking at the light peptide (low to high concentration left to right) in this data set, you can see the purple transition is too large, in comparison to the heavy peptide, and it decreases in size as the light peptide concentration increases. This is a quick way to visualize interferences in a response curve or any set of sample replicates.

## Slide 45: Standards
Are they necessary?
Peptide or Protein?
What if I don’t have any?

## Slide 46: What you might normally see without the IS
Peak Area

## Slide 47: What you should see with the IS (same data)
Peak Area Ratio (L:H)

## Slide 48: Types of Standards
Labeled Reference Peptide
  One or more heavy peptides used for normalization
Global
  Addition of heavy-labeled sample of entire proteome (SILAC)
Labeled Synthetic Tryptic Peptides
  Isotopically labeled peptides for each target
“Winged” or “Extended” Synthetic Peptides
  Contain native amino acids and are spiked pre-digest to account for digestion variability
Heavy Labeled Proteins
  Provide

## Slide 49: What you might normally see without the IS: Greater Variability
Internal Standard
Labeled Reference Peptide
Label-Free no normalization
Zhang, H., et al (2011) MCP 10(6):M110.006592
While costly, internal standards can compensate for variability in sample preparation and data acquisition, resulting in better precision.

## Slide 50: Protein Standards Closest to Ideal: Added early in workflow
Abbatiello, Schilling, Mani, et al, MCP 2015

## Slide 51
Technical and Process Variability Assessed From Digestion Controls and SIS Peptide Spikes
Light Peak Area from Protein Digestion Controls
13C/15N Peak Area from post-desalt peptide spikes
Process Variability
Technical Variability
Raw Peak Area CV assessed over ~150 injections
Abbatiello, Schilling, Mani, et al 2015 MCP

## Slide 52: Targeted Assay Optimization
What should you spend your time on?
Digestion
Denaturation
Desalting
Peptide selection
Transition selection
Chromatography
Collision Energy
IS Spike Level
Retention Time Scheduling
Dwell Time/ Injection Time
Cycle Time

## Slide 53: Key and Important Aspects to Targeted Peptide Methods
Reproducibility (Precision)
Accuracy
Sensitivity
Robustness
Dynamic Range
https://sites.google.com/a/apaches.k12.in.us/mr-evans-science-website/accuracy-vs-precision
https://en.wikipedia.org/wiki/Calibration_curve#/media/File:Calibration_curve.png

## Slide 54: Summary
Consult the Literature!
Generate data!
Review those data!
Evaluate if the data make sense!
Repeat!

## Slide 55: Introduction to Skyline
Brendan